## App V - In study validation
Accuracies of calibrators and QC samples
of Midazolam - RUN01
ID of sample | Nominal conc. [ng/mL] | Measured conc. [ng/mL] | Accuracy [%]
Calibrators
C1 | 1 | 0.842 | -15.8
C2 | 2 | 1.89 | -5.5
C3 | 5 | 5.23 | 4.6
C4 | 10 | 9.8 | -2
C5 | 20 | 21.7 | 8.5
C6 | 50 | 51.7 | 3.4
C7 | 100 | 85.5 | -14.5
C8 | 200 | 189 | -5.5
C9 | 500 | 518 | 3.6
C10 | 1000 | 1020 | 2
Quality control samples
QCH | 800 | 904 | 13
QCM | 300 | 296 | -1.33
QCL | 3 | 2.89 | -3.67
QCH | 800 | 906 | 13.2
QCM | 300 | 269 | -10.3
QCL | 3 | 3.02 | 0.667
Acceptance criteria:
Accuracy ≤ ± 20% for C1, ≤ ± 15% for all other samples
Accuracy %: (measured conc.- nominal conc.)/nominal conc. * 100
Acceptance of the run:
Accuracy of at least 75% of calibrators and 67% of QC samples
(at least 50% at each conc. level) meet the acceptance criteria; r: ≥ 0.99
To accept the results of the diluted samples at least 67% of the diluted high
concentration QC samples should be within ±15% of their respective nominal value.
Parameters of the calibration curve:
Equation (correlation):
Accuracies of calibrators and QC samples
of Midazolam - RUN02
ID of sample | Nominal conc. [ng/mL] | Measured conc. [ng/mL] | Accuracy [%]
Calibrators
C1 | 1 | 1.061 | 6.1
C2 | 2 | 1.86 | -7
C3 | 5 | 5.13 | 2.6
C4 | 10 | 9.02 | -9.8
C5 | 20 | 23.9 | 19.5
C6 | 50 | 55.5 | 11
C7 | 100 | 107.3 | 7.3
C8 | 200 | 180 | -10
C9 | 500 | 493 | -1.4
C10 | 1000 | 1070 | 7
Quality control samples
QCH | 800 | 756 | -5.5
QCM | 300 | 312 | 4
QCL | 3 | 3.19 | 6.33
QCH | 800 | 755 | -5.62
QCM | 300 | 283 | -5.67
QCL | 3 | 2.71 | -9.67
Dilution samples
dil-QCH-15x | 12000 | 11000 | -8.33
dil-QCH-15x | 12000 | 10700 | -10.8
dil-QCH-15x | 12000 | 11800 | -1.67
Acceptance criteria:
Accuracy ≤ ± 20% for C1, ≤ ± 15% for all other samples
Accuracy %: (measured conc.- nominal conc.)/nominal conc. * 100
Acceptance of the run:
Accuracy of at least 75% of calibrators and 67% of QC samples
(at least 50% at each conc. level) meet the acceptance criteria; r: ≥ 0.99
To accept the results of the diluted samples at least 67% of the diluted high
concentration QC samples should be within ±15% of their respective nominal value.
Parameters of the calibration curve:
Equation (correlation):
Accuracies of calibrators and QC samples
of Midazolam - RUN03
ID of sample | Nominal conc. [ng/mL] | Measured conc. [ng/mL] | Accuracy [%]
Calibrators
C1 | 1 | 0.991 | -0.9
C2 | 2 | 2.19 | 9.5
C3 | 5 | 5.24 | 4.8
C4 | 10 | 8.79 | -12.1
C5 | 20 | 20 | 0
C6 | 50 | 46 | -8
C7 | 100 | 87.3 | -12.7
C8 | 200 | 225 | 12.5
C9 | 500 | 528 | 5.6
C10 | 1000 | 937 | -6.3
Quality control samples
QCH | 800 | 893 | 11.6
QCM | 300 | 276 | -8
QCL | 3 | 3.62 | 20.7
QCH | 800 | 755 | -5.62
QCM | 300 | 274 | -8.67
QCL | 3 | 2.73 | -9
Dilution samples
dil-QCH-15x | 12000 | 11100 | -7.5
dil-QCH-15x | 12000 | 10800 | -10
dil-QCH-15x | 12000 | 13300 | 10.8
Acceptance criteria:
Accuracy ≤ ± 20% for C1, ≤ ± 15% for all other samples
Accuracy %: (measured conc.- nominal conc.)/nominal conc. * 100
Acceptance of the run:
Accuracy of at least 75% of calibrators and 67% of QC samples
(at least 50% at each conc. level) meet the acceptance criteria; r: ≥ 0.99
To accept the results of the diluted samples at least 67% of the diluted high
concentration QC samples should be within ±15% of their respective nominal value.
Parameters of the calibration curve:
Equation (correlation):
Accuracies of calibrators and QC samples
of Midazolam - RUN04
ID of sample | Nominal conc. [ng/mL] | Measured conc. [ng/mL] | Accuracy [%]
Calibrators
C1 | 1 | 0.955 | -4.5
C2 | 2 | 2.04 | 2
C3 | 5 | 5.62 | 12.4
C4 | 10 | 9.4 | -6
C5 | 20 | 17.9 | -10.5
C6 | 50 | 50.1 | 0.2
C7 | 100 | 91.4 | -8.6
C8 | 200 | 188 | -6
C9 | 500 | 469 | -6.2
C10 | 1000 | 920 | -8
Quality control samples
QCH | 800 | 858 | 7.25
QCM | 300 | 277 | -7.67
QCL | 3 | 2.72 | -9.33
QCH | 800 | 779 | -2.62
QCM | 300 | 285 | -5
QCL | 3 | 2.7 | -10
Dilution samples
dil-QCH-15x | 12000 | 10800 | -10
dil-QCH-15x | 12000 | 12700 | 5.83
dil-QCH-15x | 12000 | 12600 | 5
Acceptance criteria:
Accuracy ≤ ± 20% for C1, ≤ ± 15% for all other samples
Accuracy %: (measured conc.- nominal conc.)/nominal conc. * 100
Acceptance of the run:
Accuracy of at least 75% of calibrators and 67% of QC samples
(at least 50% at each conc. level) meet the acceptance criteria; r: ≥ 0.99
To accept the results of the diluted samples at least 67% of the diluted high
concentration QC samples should be within ±15% of their respective nominal value.
Parameters of the calibration curve:
Equation (correlation):

## App VI - Between-run
Results of QC samples: Between-run calculated mean concentrations, accuracies and precisions
Midazolam
Between-run accuracy and precision of quality control samples
Sample name | Nominal conc. [ng/mL] | Accuracy % | Conc. [ng/mL] | Precision %
QCH | 800 | 3.21 | 826 | 8.59
QCL | 3 | -1.75 | 2.95 | 11
QCM | 300 | -5.33 | 284 | 4.92
dil-QCH-15x | 12000 | -2.96 | 11600 | 8.49
Acceptance criteria: The between-run accuracy should be within
±15 % of the nominal one and the CV% should be ≤15 %.

## App_conc
| Group | Gp1M |  |  |  |  | Gp1F | [5 empty] | Mean | CV%
| Animal | An 1001 | An 1002 | An 1003 | An 1004 | An 1005 | An 1501 | An 1502 | An 1503 | An 1504 | An 1505 | An 1506
Day | Time | Gp1M An 1001: Concentration [ng/mL]
Day 1 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 1h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 4h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 6h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 8h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 12h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 24h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Day 28 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 1h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 4h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 6h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 8h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 12h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 24h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Group | Gp1M An 1001: Gp2M; Gp1M An 1005: Gp2F; Gp1F An 1504: Mean; Gp1F An 1505: CV%
| Animal | An 2001 | An 2002 | An 2003 | An 2004 | An 2501 | An 2502 | An 2503 | An 2504
Day | Time | Gp2M An 2001: Concentration [ng/mL]
Day 1 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 162 | 172 | 174 | 198 | 204 | 206 | 162 | 209 | 186 | 10.9
| 1h | 881 | 924 | 927 | 913 | 832 | 922 | 984 | 915 | 912 | 4.73
| 2h | 1140 | 1400 | 1280 | 1390 | 1260 | 1110 | 1280 | 1360 | 1280 | 8.46
| 4h | 942 | 1020 | 712 | 1160 | 982 | 817 | 620 | 1090 | 918 | 20.4
| 6h | 484 | 715 | 389 | 599 | 567 | 372 | 338 | 686 | 519 | 28
| 8h | 251 | 366 | 198 | 486 | 327 | 199 | 153 | 421 | 300 | 39.6
| 12h | 61.1 | 130 | 42.9 | 160 | 106 | 59 | 30.9 | 183 | 96.5 | 58.8
| 24h | 1.8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.225 | 283
Day 28 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 181 | 191 | 188 | 161 | 158 | 162 | 182 | 195 | 177 | 8.32
| 1h | 805 | 791 | 835 | 992 | 948 | 780 | 902 | 1020 | 884 | 10.7
| 2h | 1090 | 1290 | 1050 | 1240 | 1100 | 970 | 891 | 1420 | 1130 | 15.4
| 4h | 558 | 662 | 479 | 704 | 613 | 584 | 368 | 1060 | 629 | 32.6
| 6h | 240 | 419 | 168 | 337 | 378 | 242 | 86.5 | 653 | 315 | 55.5
| 8h | 89.3 | 188 | 72 | 168 | 156 | 81.6 | 29.7 | 326 | 139 | 67.1
| 12h | 14.4 | 45.6 | 9.59 | 32.4 | 28.8 | 13.8 | 2.08 | 129 | 34.5 | 118
| 24h | 0 | 0 | 0 | 4.25 | 0 | 0 | 3.89 | 0 | 1.02 | 185
| Group | Gp2M An 2001: Gp3M; Gp2F An 2501: Gp3F | Mean | CV%
| Animal | An 3001 | An 3002 | An 3003 | An 3004 | An 3501 | An 3502 | An 3503 | An 3504
Day | Time | Gp3M An 3001: Concentration [ng/mL]
Day 1 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 7.56 | 333 | 265 | 359 | 341 | 297 | 294 | 360 | 282 | 41.1
| 1h | 1680 | 1540 | 1420 | 1700 | 1480 | 1310 | 1290 | 1710 | 1520 | 11.2
| 2h | 2180 | 2250 | 2020 | 2110 | 2670 | 2250 | 1920 | 2630 | 2250 | 11.9
| 4h | 1230 | 1800 | 1420 | 1870 | 1460 | 1240 | 1210 | 1620 | 1480 | 17.5
| 6h | 868 | 1110 | 602 | 1070 | 930 | 653 | 586 | 1380 | 900 | 31.3
| 8h | 415 | 732 | 338 | 823 | 503 | 340 | 240 | 794 | 523 | 43.8
| 12h | 132 | 245 | 72.9 | 326 | 184 | 105 | 50.8 | 337 | 182 | 61.1
| 24h | 3.33 | 5.31 | 0 | 2.23 | 3.49 | 3.23 | 0 | 0 | 2.2 | 91.3
Day 28 | 0h | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 0.5h | 331 | 345 | 337 | 339 | 305 | 281 | 255 | 289 | 310 | 10.6
| 1h | 1270 | 1470 | 1420 | 1540 | 1660 | 1280 | 1410 | 1500 | 1440 | 9.02
| 2h | 1880 | 2250 | 1950 | 2140 | 2340 | 2230 | 1480 | 2390 | 2080 | 14.5
| 4h | 1030 | 1420 | 919 | 1340 | 1340 | 1020 | 564 | 1540 | 1150 | 28.1
| 6h | 338 | 705 | 353 | 513 | 646 | 424 | 185 | 940 | 513 | 47.1
| 8h | 143 | 306 | 110 | 246 | 285 | 154 | 52.6 | 677 | 247 | 79
| 12h | 20.4 | 90 | 16.8 | 56.9 | 63.4 | 20.6 | 2.82 | 230 | 62.6 | 118
| 24h | 0 | 50.2 | 0 | 0 | 2.5 | 0 | 135 | 30.5 | 27.3 | 174